EMpower and Firmware v2.56.1 Release Notes
August 4 2026
=======================================
	Operating systems supported
=======================================
Windows
        * Permissions required: Must be logged in as Administrator
        * Versions tested:
                - Windows 10, 64-bit
                - Windows 11, 64-bit
        * NOTE: Windows 10S and Windows 11S are not supported
Linux
        * Permissions required: Standard user
        * Versions tested:
                - Ubuntu 20.04.4
                - Ubuntu 24.04.4
Mac
        * Permissions required: Standard User
        * Versions tested:
                - macOS 10.15.4, Catalina
                - macOS 12.0 (Monterey - Apple Silicon)
                - macOS  14.4 (Sonoma)
=======================================
What's new in the Latest Version
=======================================

== System Enhancements ==

●	Added OS upgrade support for RMT08-B receivers.
●	Corrected phase inconsistencies in certain stacked CSAMT frequencies. All stacks at the same frequency are now phase-consistent.
●	This release increases the stability of measurement channels when they operate close to the maximum temperature limit of the instrument.

== EMpower New Features ==
●	New automated calculator for TDEM ramp time, based on transmitter loop layout.
●	CSAMT layout grids are now referred to true north. The user can now indicate the magnetic declination for the grid.
●	Added a new user-adjustable manual gate delay field for TDEM recordings.
●	Added a Go to Recording action to the context menu for TDEM processed sites and workbenches. This action opens the Recording Library, selects the associated TDEM recording on the map and timeline, and displays its recording information.
●	Added an Advanced Settings button to the Recording Information dialog for generating higher-density TDEM time gates.

== EMpower Bug Fixes ==
●	Fixed an issue in the TDEM stack editor where the decay plot displayed incorrect data points at early times.
●	Fixed an issue where the zoom level in the recording library map would reset when saving in the CSAMT station editor. 
●	Fixed an issue in the processing queue dialog plots where the frequency range was not limited based on the selected sensor type.
●	Removed the (undesired) capability of exporting Time Domain recordings as a CSV  files containing frequency domain calculated fields.